# Califia Perfusion Simulator

Getting Started with Califia Perfusion Simulator

January 2023

# Important Information

No part of this work may be reproduced in any form or by any means - graphic, electronic, or mechanical, including photocopying, recording, taping, or information storage and retrieval systems - without the written permission of Biomed Simulation Inc.

Products that are referred to in this document may be either trademarks and/or registered trademarks of the respective owners. The publisher and the author make no claim to these trademarks.

While every precaution has been taken in the preparation of this document, the publisher and the author assume no responsibility for errors or omissions, or for damages resulting from the use of information contained in this document or from the use of programs and source code that may accompany it. In no event shall the publisher and the author be liable for any loss of profit or any other commercial damage caused or alleged to have been caused directly or indirectly by this document.

Under no circumstances should this simulator system be used to guide the management of any living patient.

Califia® is a registered trademark of Biomed Simulation Inc.

US Patent No.	            	14051442

European Patent No.	2906107

Canadian Patent No.	2887344

Singapore Patent No. 	11201502855X

Chinese Patent No. 	ZL 201380062871X

US Patent Pending 	PCT/US2018/012009

Califia® is in conformity with the following standards

EVS-EN 61326-1:2013	Electrical equipment for measurement, control and laboratory use – EMC requirements – Part 1: General requirements.

EVS-EN 61010-1:2010 	Safety requirements for electrical equipment for measurement, control and laboratory use – Part 1: General Requirements

Image Source: www.biomedsimulation.com

Copyright © Biomed Simulation Inc. 2011-2023. All Rights Reserved.

# Contents

Index

Important Information	2

Contents	3

1. Introduction to the Califia Perfusion Simulator	8

1.1 Components of the Califia Perfusion Simulator	8

Figure 1.1. Califia Perfusion Simulator Setup	8

Figure 1.2. CPM Hydraulic Ports	9

1.2 CPM connections for Cardiopulmonary Bypass (CPB)	10

Table 1.1 HLM connections to CPM	10

Figure 1.3 CPB connections to CPM	10

1.3 CPM connections for ECMO	11

1.3.1 VA ECMO	11

Table 1.2 VA ECMO connections to CPM	11

1.3.2 VV ECMO	12

Table 1.3 VV ECMO connections to CPM	12

Figure 1.5  VV ECMO with CPM	12

1.3.3. VAV ECMO	13

Table 1.4 VAV ECMO connections with CPM	13

Figure 1.6. VAV ECMO with CPM	13

2. A brief tour of the Califia Simulator software	14

Figure 2.2. Califia Simulator Opening Screen	14

2.1 Modes of Simulation: SIM and CPM	15

2.2 Modes of Operation: CPB and ECMO	15

2.3 Clinical Scenario files	15

2.4 Instructor’s Panel	16

Figure 2.3  Instructor's Panels for CPB and ECMO respectively	16

2.5 Learner Screen	17

Figure 2.4 Learner Screen for CPB and ECMO respectively	17

2.6 The Califia 3D ICU environment (ECMO only)	18

Figure 2.5 Califia 3D ICU (ECMO only)	18

3. Prepare the Califia Patient Module	19

3.1 Priming CPM	19

Figure 3.1 CPM with tubing clamped	19

Figure 3.2. Priming CPM	20

3.2 Return to the Opening Screen	20

4. Let’s go on pump! – a CPB simulation	21

4.1 SIM vs Califia Patient Module	21

4.2 Use a Standardized Patient scenario	22

Figure 4.1. Opening Screen » CPB mode	22

4.3 Patient Demographics	23

Figure 4.2. Patient demographics in Instructor's Panel	23

4.4 CVP Gain	24

Table 4.1. BSA vs CVP Gain	24

Figure 4.3. Califia Calibration Panel > CVP Gain	24

4.5 CVP	25

Figure 4.4. Fill CPM with water to match target CVP	25

Figure 4.5. Set Target CVP in SIM mode	26

4.6 Arterial Pressure	27

Figure 4.6. Set Systolic and Diastolic Arterial Pressure	27

4.7 Blood Gases	28

Figure 4.7. Set Blood Gases	28

4.8 Select Delivery and Drainage Cannulas	29

Figure 4.8. Cannula selection	29

4.9 Set Gas Blender	30

Figure 4.9 Set Gas Blender	30

4.10 Let’s go on pump!	31

Figure 4.10 Arterial and CVP decreasing	31

Figure 4.11 Prime circuit using Fluids Panel	32

Figure 4.12 Starting bypass (SIM)	33

4.11 Give Cardioplegia	34

Figure 4.13 Prepare for Cardioplegia delivery	34

Figure 4.14 Heart ECG under Cardioplegia delivery	35

Figure 4.15. Heart ECG under Cardioplegia delivery (SIM)	36

4.12 Event: Cannula malposition	37

Figure 4.16. Arterial Cannula set to 25% opened	37

Figure 4.17. High line pressure	38

Figure 4.18. High arterial line pressures (SIM)	39

4.13 Remove Cross Clamp	40

Figure 4.19. Open Cross Clamp	40

Figure 4.20. Open Cross Clamp (SIM)	41

4.14 Come off pump!	42

Figure 4.21  Wean off pump	42

Figure 4.22. Wean off pump (SIM)	43

4.15 Return to Opening Screen	44

Figure 4.23. End Simulation	44

4.16 Create a CPB Simulation Scenario File	45

4.16.1 Access Scenario Editor -	45

Figure 4.24. Access MENU in Opening Screen then Scenario Editor	45

4.16.2 Save Scenario File -	46

Figure 4.25. Save As... in Scenario Editor	46

4.16.3 New scenario file in Opening Screen -	47

Figure 4.26. Opening Screen - New Scenario File	47

4.16.4 Scenario File Step 2: Select cannulas -	48

Figure 4.27. Step 2 edits	48

4.16.5 Scenario File Step 3: Go on pump -	49

Figure 4.28 Step 3 edits	49

4.16.6 Scenario File Step 4: Give Cardioplegia -	50

Figure 4.29 Step 4 edits	50

4.16.7 Scenario File Step 5: Event: Cannula malposition -	51

Figure 4.30 Step 5 edits	51

4.16.8 Scenario File Step 6: Open Cross Clamp	52

Figure 4.31. Step 6 edits	52

4.16.9 Scenario File Step 7: Come off pump -	53

5. Let’s go on ECMO! – an ECMO simulation	54

5.1 SIM vs Califia Patient Module	54

5.2 Use a Standardized Patient scenario	55

5.3 Patient Demographics	56

Figure 5.2. Patient demographics in Instructor's Panel	56

5.4 CVP Gain	57

Table 5.1. BSA vs CVP Gain	57

Figure 5.3. Califia Calibration Panel > CVP Gain	57

5.5 CVP	58

Figure 5.4. Fill CPM with water to match target CVP	58

Figure 5.5. Set Target CVP in SIM mode	59

5.6 Arterial Pressure	60

Figure 5.6. Set Systolic and Diastolic Arterial Pressure	60

5.7 Ventilator Settings	61

Figure 5.7. Ventilator settings	61

5.8 Blood Gases	62

Figure ‎5.8. Set Blood Gases	62

5.9 Select Delivery and Drainage Cannulas	63

Figure 5.9 . Cannula selection	63

5.10 Set Gas Blender	64

Figure 5.10. Set Gas Blender	64

5.11 Let’s go on pump!	65

Figure 5.11. Arterial and CVP decreasing	65

Figure 5.12. Prime circuit using Fluids Panel	66

Figure 5.13. Starting bypass (SIM)	67

5.12 Give Cardioplegia	68

Figure 5.14. Prepare for Cardioplegia delivery	68

Figure 5.15 Heart ECG under Cardioplegia delivery	69

Figure 5.16. Heart ECG under Cardioplegia delivery (SIM)	70

5.13 Event: Cannula malposition	71

Figure 5.17. Arterial Cannula set to 25% opened	71

Figure 5.18. Line pressure increase (line manometer photo)	72

Figure 5.19. High arterial line pressures (SIM)	73

5.14 Remove Cross Clamp	74

Figure 5.20. Open Cross Clamp	74

Figure 5.21. Open Cross Clamp (SIM)	75

5.15 Come off pump!	76

Figure 5.22. Wean off pump	76

Figure 5.23. Wean off pump (SIM)	77

5.16 Return to Opening Screen	78

Figure 5.24. End Simulation	78

5.17 Create a CPB Simulation Scenario File	79

5.18 Access Scenario Editor	79

Figure 5.25. Access MENU in Opening Screen then Scenario Editor	79

5.19 Save Scenario File	80

Figure 5.26. Save As... in Scenario Editor	80

5.20 New scenario file in Opening Screen	81

Figure 5.27. Opening Screen - New Scenario File	81

5.21 Scenario File Step 2: Select cannulas	82

Figure 5.28. Step 2 edits	82

5.22 Scenario File Step 3: Go on pump	83

Figure 5.29. Step 3 edits	83

5.23 Scenario File Step 4: Give Cardioplegia	84

Figure 5.30. Step 4 edits	84

5.24 Scenario File Step 5: Event: Cannula malposition	85

Figure 5.31. Step 5 edits	85

5.25 Scenario File Step 6: Open Cross Clamp	86

Figure 5.32. Step 6 edits	86

5.26 Scenario File Step 7: Come off pump	87

Figure 5.33. Step 7 edits	87

#### 1

Introduction to the Califia Perfusion Simulator

The Califia Perfusion Simulator system can simulate a patient before, during and after CardioPulmonary Bypass (CPB) for open heart surgery or long-term support with ExtraCorporeal Membrane Oxygenation (ECMO).

## Components of the Califia Perfusion Simulator

The three components of the Califia Perfusion Simulator, shown in Figure 1.1, are the Califia Patient Module (CPM), the Laptop Computer and the Touchscreen Monitor. Apply power to each component.

The CPM connects to the Laptop Computer using an ethernet cable. Two cables connect the Touchscreen Monitor to the laptop, an HDMI cable and a USB cable.

### Figure 1.1. Califia Perfusion Simulator Setup

The CPM hydraulic ports, ‎Figure 1.2, connect to either a Heart Lung Machine (HLM), in the case of CPB, or an ECMO system.

| CPM hydraulic ports | Description |
| --- | --- |
| Arterial | Main delivery port into the CPM. It connects to the oxygenator outflow line. |
| Venous | The CPM’s drainage port. It connects to a venous reservoir in the case of CPB or directly to the pump’s inlet port in ECMO circuit configurations. |
| Cardioplegia | Secondary deliver port into the CPM. Used for Cardioplegia delivery during CPB or as a second delivery port in hybrid ECMO circuit configurations. |
| Vent | Secondary drainage port. During CPB, it can connect to the Left Ventricle Vent pump. |
| Drain | Tertiary drainage port. It is used to drain the CPM’s reservoir during the teardown process. |
| Overflow | It provides runoff should the CPM’s reservoir overfill. |

### Figure 1.2. CPM Hydraulic Ports

## CPM connections for Cardiopulmonary Bypass (CPB)

Figure 1.3 illustrates connections from a simplified CPB circuit to the CPM. Table 1.1 summarizes these connections.

| HLM | CPM |
| --- | --- |
| Systemic blood pump | Arterial |
| Venous Reservoir | Venous |
| Cardioplegia delivery | Cardioplegia |
| LV Vent pump | Vent |

### Table 1.1 HLM connections to CPM

### Figure 1.3 CPB connections to CPM

Note - If the Vent port is not used, connect a ¼” tubing segment and clamp it.

## CPM connections for ECMO

    - VA ECMO  - Figure 1.4 illustrates connections between a VA ECMO circuit and the CPM.  Table 1.2 summarizes these connections.

| ECMO | CPM | NOTES |
| --- | --- | --- |
| Oxygenator output port | Arterial |  |
| ECMO Pump inlet | Venous |  |
|  | Cardioplegia | Connect these together |
|  | Vent |  |

### Table 1.2 VA ECMO connections to CPM

Figure 1.4 VA ECMO with CPM

    - VV ECMO --Figure 1.5, illustrates connections between a VV ECMO circuit and the CPM. Table 1.3 summarizes these connections.

| ECMO | CPM | NOTES |
| --- | --- | --- |
| Oxygenator output port | Arterial |  |
| ECMO Pump inlet | Venous |  |
|  | Cardioplegia | Connect these together |
|  | Vent |  |

### Table 1.3 VV ECMO connections to CPM

### Figure 1.5  VV ECMO with CPM

    - VAV ECMO -Figure 1.6 illustrates connections between a VAV ECMO circuit and the CPM. Table 1.4 summarizes these connections.

| ECMO | CPM | NOTES |
| --- | --- | --- |
| Delivery main branch | Arterial |  |
| ECMO Pump inlet | Venous |  |
| Delivery IJ branch | Cardioplegia |  |
|  | Vent | Clamped tubing |

### Table 1.4 VAV ECMO connections with CPM

### Figure 1.6. VAV ECMO with CPM

#### 2

# A brief tour of the Califia Simulator software

To launch the Califia Simulator software, double-click the Califia Simulator icon located in the laptop computer’s desktop. Refer to Figure 2.1.

Figure 2.1. Califia Simulator desktop icon

### Figure 2.2. Califia Simulator Opening Screen

## Modes of Simulation: SIM and CPM

Refer to ① in  Figure 2.2.

SIM 	this mode does not use the Califia Patient Module. All relevant equipment, Heart Lung Machine (HLM) in CPB or ECMO circuit in ECMO, is virtualized and made accessible in the software.

Califia Patient Module 	this mode uses the CPM and real life support equipment such as a HLM or ECMO circuit. If the Califia Simulator software detects the CPM present, this selection becomes available.

## Modes of Operation: CPB and ECMO

Refer to ② in Figure 2.2

Select either CPB or ECMO to step into the selected simulation application. The next step is to select a clinical simulation scenario then press the Continue button on the lower right of the window.

## Clinical Scenario files

Refer to ③ in  Figure 2.2.

Many of the buttons in the Clinical Scenarios section provide easy links to scenario files. For this Getting Started process, we will use one of the Standardized Patient scenarios.

For CPB, select the button labeled SP CPB Adult.

For ECMO, select the button labeled SP VA ECMO Adult.

Either of these simulation scenarios will configure the patient as an adult in need of either CPB or VA ECMO support respectively.

## Instructor’s Panel

Figure 2.3 shows the Instructor’s Panel for CPB and ECMO side-by-side.

These panels have a similar layout. The CPB Instructor’s Panel shows its space divided into sections.

① refers to common numerical indicators and entry fields such as blood gas sampling values.

② is the scenario file and steps section.

③ shows the grouping of text fields.

④ displays a different top half based on the button selection.

### Figure 2.3  Instructor's Panels for CPB and ECMO respectively

## Learner Screen

Figure 2.4 shows the Learner Screen for CPB and ECMO side-by-side.

As in the case of the Instructor’s Panels, these screens have a similar layout. Refer to the numbers in the image for the following descriptions

① displays the patient’s vital signs in the form of traces and numerically

② is an area used to display relevant medical device panels like inline blood gas analysers, heater-coolers and clotting timers.

③ is a blood sampling report. This is a floating panel that can be moved to different locations. There are other similar floating panels available for either CPB or ECMO.

④ refers to the gas blender elements: Sweep and Oxygenator FiO2.

⑤ is the bottom ribbon with access to panels and options such as Alarms and Media panels.

⑥ is shown in the ECMO learner screen and refers to the Ventilator settings area.

### Figure 2.4 Learner Screen for CPB and ECMO respectively

## The Califia 3D ICU environment (ECMO only)

Figure 2.5 shows the 3D Intensive Care Unit (ICU), available for ECMO simulation. The learner can navigate the environment and interact with the medical devices in the space.

### Figure 2.5 Califia 3D ICU (ECMO only)

#### 3

# Prepare the Califia Patient Module

## Priming CPM

Prepare the CPM with tubing connections, refer to Figure 3.1

- The Arterial and Venous ports attach to ½” tubing.

- The Cardioplegia and Vent ports attach to ¼” tubing.

- Make each tubing segment at least 6” in length. Clamp all 4 tubings to avoid leaking water.

### Figure 3.1 CPM with tubing clamped

Power ON the CPM, wait about 30 seconds and select the Califia Patient Module simulation mode in the Opening Screen, then press Continue to step into either CPB or ECMO per your preference.

Refer to Figure 3.2 to prime the CPM.

① Use a pitcher to carefully pour about 2.5 Liters of water into the CPM’s reservoir.

  - Notice patient’s vitals such as arterial pressure and central venous pressure increase in value; in essence, the vitals correspond to those of a living patient.

  - Ensure all tubing is clamped to avoid any water leak

② Open the side door to align the silicon tubing in peristaltic pump’s raceway

③ Do not operate peristaltic pump while adjusting tubing

④ Close peristaltic pump lid

⑤ Close side door

⑥ In the Instructor’s Panel » Califia Controls section, press the Pressurize button to operate the peristaltic pump and prime the CPM. After about 5 minutes of operation, press the Pressurizing button to stop the pump, the CPM should be primed.

Figure 3.2. Priming CPM

High flow rates at lower pressure are achieved with the CPM placed 18” to 20” above venous reservoir (CPB) or pump inlet (ECMO).

## Return to the Opening Screen

If you wish to choose a different simulation scenario as instructed in one of the next two next sections, press the End Simulation button in the Instructor’s Panel to return to the Opening Screen.

#### 4

# Let’s go on pump! – a CPB simulation

This chapter steps through a CPB simulation. You can follow along.

Approximately 60 minutes to complete this chapter.

## SIM vs Califia Patient Module

Refer to Figure 4.1, ①.

|  | Sections in this chapter headlined with this image refer to software-only simulation. No CPM or real pump is involved. |
| --- | --- |
|  | Sections in this chapter headlined with this image use a Heart Lung Machine connected to the CPM. |

## Use a Standardized Patient scenario

Select one of the three SP CPB scenarios available, Figure 4.1, ②:

- SP CPB Adult

- SP CPB Infant

- SP CPB Pediatric

Press Continue button, ③, to step into the simulation.

### Figure 4.1. Opening Screen » CPB mode

## Patient Demographics

Refer to Figure 4.2  review and change the following:

① Age

② HCT

③ Gender

④ Weight

⑤ Height

### Figure 4.2. Patient demographics in Instructor's Panel

## CVP Gain

In Instructor’s Panel, press the Calibration Panel button under Califia Controls to bring up the Califia Calibration Panel, Figure 4.3. Set CVP Gain as dictated in Table 4.1

| BSA (m2) | CVP Gain |
| --- | --- |
| < 1 | 3.5 |
| Between 1 & 2 | 3.0 |
| > 2 | 2.5 |

### Table 4.1. BSA vs CVP Gain

### Figure 4.3. Califia Calibration Panel > CVP Gain

## CVP

Set CVP to 10 mmHg

Fill Califia’s Reservoir until CVP is 10 mmHg, ① in  Figure 4.4

### Figure 4.4. Fill CPM with water to match target CVP

Refer to Figure 4.5, go to Instructor’s Panel » Califia Controls

① Enter 10 for Target CVP then press <Enter>

② Volume in the virtual Califia’s Reservoir changes until CVP reaches 10, ③.

### Figure 4.5. Set Target CVP in SIM mode

## Arterial Pressure

Refer to Figure 4.6, in Instructor’s Panel » Cardiovascular Vars,

① Press Set Arterial Pressure button.

② Specify desired Systolic and Diastolic Arterial Pressures.

③ Press Target button.

④ It takes a few seconds to find values for LV Contractility and SVR.

⑤ Press Save Step to save all changes to scenario file.

Close Model Settings panel once targetting is complete.

### Figure 4.6. Set Systolic and Diastolic Arterial Pressure

## Blood Gases

Refer to Figure 4.7, go to Instructor’s Panel » Blood Chemistry.

① Press Calculate Desired Gases button

② Specify desired pO2, pCO2 and pH for arterial blood

③ Press Target button

④ It takes several minutes to find values for HCO3, Dead Space and Pulmonary Shunt Fraction, the last two parameters are in the Pulmonary Vars section.

⑤ Press Save Step to save all changes to scenario file.

Close Model Settings panel once targetting is complete.

### Figure 4.7. Set Blood Gases

## Select Delivery and Drainage Cannulas

Refer to Figure 4.8 to select Delivery (arterial) and Drainage (venous) Cannulas. In Instructor’s Panel » Califia Controls,

① Press Cannula Editor button.

②a In Cannula Editor, select a Delivery Cannula from drop-down list then press Apply button.

③a Cannula label and corresponding valve opening is shown.

②b Change to Drainage Cannula, select cannula from drop-down list then press Apply button.

③b Cannula label and corresponding valve opening is shown.

Exit Cannula Editor

Pressure-Flow profiles for cannulas are modeled using respective electromechanical valves in CPM. The same performance profiles are also used in SIM mode.

### Figure 4.8. Cannula selection

## Set Gas Blender

Refer to Figure 4.9, gas blender is set in the learner screen.

① Bring up bottom ribbon and select Oxygenator Controls button

② Set Gas Sweep Rate

③ Set Oxygenator FiO2 by pressing green button then turning blue knob

### Figure 4.9 Set Gas Blender

## Let’s go on pump!

Remove clamp on venous drainage line to drain patient while flowing using systemic pump.

Notice Arterial Pressure and CVP values decreasing as patient transitions support from native heart to heart lung machine, Figure 4.10.

### Figure 4.10 Arterial and CVP decreasing

Prime virtual circuit, refer to Figure 4.11

Bring up bottom ribbon

① Press Fluids Panel button

② Press HLM Controls button

In Fluids Panel

③a Select Crystalloid from Fluids drop-down menu

③b Enter 1000 mL of Volume

③c Light up the Priming button

③d Press Apply button

④ Notice Venous Reservoir fills up with this priming volume

### Figure 4.11 Prime circuit using Fluids Panel

To go on bypass, refer to Figure 4.12,

① Gradually open Venous Occluder

② Start increasing systemic pump flow

Virtual systemic pump is a roller pump

③ Watch for arterial pressure waveform flattening though as flow is increased, mean pressure rises

### Figure 4.12 Starting bypass (SIM)

## Give Cardioplegia

Refer to Figure 4.13, in Instructor’s Panel » Cardioplegia,

① Set slider to Ante (Antegrade Cardioplegia delivery)

② Close X Clamp

### Figure 4.13 Prepare for Cardioplegia delivery

Flow into Cardioplegia port. As flow is registered in the simulation software, the heart ECG will step through various patterns until the heart is completely arrested, Figure 4.14

### Figure 4.14 Heart ECG under Cardioplegia delivery

Refer to Figure 4.15:

① Close X Clamp in HLM Controls panel

② Increase Cardioplegia pump flow

③ Notice ECG pattern change

### Figure 4.15. Heart ECG under Cardioplegia delivery (SIM)

## Event: Cannula malposition

Refer to Figure 4.16, in Instructor’s Panel » Califia Controls,

① Set Delivery Cannula (%) slider to about 25%

② Arterial line pressure increases

cc

### Figure 4.16. Arterial Cannula set to 25% opened

Arterial line pressure should report a high value due to high resistance generated by the smaller arterial valve opening, Figure 4.17.

### Figure 4.17. High line pressure

Main pump type

Roller pump

High line pressure is expected until the pump’s speed is lowered.

Centrifugal pump

High line pressure should be experienced momentarily, flow rate slows down and pressure decreases accordingly.

Due to the high resistance generated by the smaller arterial valve opening, the arterial line pressure displayed using the pressure panel is very high, Figure 4.18

Line pressure (post-Oxygenator): 168 mmHg

Pre-Oxygenator pressure: 178 mmHg

### Figure 4.18. High arterial line pressures (SIM)

Set Delivery Cannula (%) back to 100% once the event is resolved.

## Remove Cross Clamp

Refer to Figure 4.19

① Open Cross Clamp in Instructor’s Panel » Cardioplegia

② Notice ECG pattern return to activity, ending in SR.

### Figure 4.19. Open Cross Clamp

Refer to Figure 4.20

① Open Cross Clamp in HLM Controls panel

② Notice ECG pattern return to activity, ending in SR.

### Figure 4.20. Open Cross Clamp (SIM)

## Come off pump!

Refer to Figure 4.21

Start weaning by partially occluding venous drain and filling patient.

① Notice patient’s heart starts to eject and venous pressure increases

Come-off systemic pump as patient’s arterial pressure returns to normal and CVP meets target value.

② Heart has taken over full body perfusion and CVP is normal value

### Figure 4.21  Wean off pump

Refer to Figure 4.22.

① Start weaning by partially occluding venous drain and filling patient.

② Notice patient’s heart starts to eject and venous pressure increases

③ Come-off systemic pump as patient’s arterial pressure returns to normal and CVP meets target value.

④ Heart has taken over full body perfusion and CVP is normal value

### Figure 4.22. Wean off pump (SIM)

## Return to Opening Screen

Refer to Figure 4.23, press End Simulation button, confirm request to return to Opening Screen.

### Figure 4.23. End Simulation

## Create a CPB Simulation Scenario File

In this section, we will repeat the CPB case presented in this chapter however we will save all stages into a scenario file.

### Access Scenario Editor -

Refer to Figure 4.24, Press MENU button, ①, followed by Scenario Editor button, ②.

### Figure 4.24. Access MENU in Opening Screen then Scenario Editor

### Save Scenario File -

Refer to Figure 4.25, in Scenario Editor,

① Press File menu selection

② Select Save As….

③ Enter a new scenario file name such as Getting Started – CPB Scenario

This name is used to create a folder which will host the scenario file and ventilator.ini file

[Optional] – the Button Selector defaults to Other however a different button can be programmed to be associated with the new scenario file

④ Press the Create button to complete the new scenario file creation

Notice the Scenario Editor title reflects the new folder location and file name

⑤ Once new scenario file and folder created, return to File menu and press Close

Press Close button to close the MENU bar, this step is not shown.

### Figure 4.25. Save As... in Scenario Editor

### New scenario file in Opening Screen -

Refer to Figure 4.26, Opening Screen

① Location of Scenario File points to the newly created Scenario File

② The selected button is Other as specified in the Save As… dialog

③ Press Continue button to load new scenario file

### Figure 4.26. Opening Screen - New Scenario File

### Scenario File Step 2: Select cannulas -

The original scenario file was saved with patient’s blood gases configured as desired, Section 4.7. [Blood Gases]

Next is to select cannulas. Follow these steps in Instructor’s Panel, refer to Figure 4.27

① Press Create New Step

② Step #2 is added to scenario table; type Select cannulas for Scenario Step then press <Enter> on keyboard or click away with the mouse

3  Follow steps in Section 4.8 to select cannulas

④ Press Save Step** button

** refers to the step having been modified and needs to be saved

⑤ In the Confirm Scenario File write dialog, select both Drainage Cannula and Delivery Cannula 1

Hold <Ctrl> key to select multiple entries

⑥ Press Confirm button to save step changes to scenario file

### Figure 4.27. Step 2 edits

### Scenario File Step 3: Go on pump -

Follow these steps, refer to Figure 4.28

① Press Create New Step

② Type Let’s go on pump for Scenario Step then press <Enter> on keyboard or click away with the mouse

③ Add following note in Skills field: Set gas blender and go on pump

④ Add following in Surgical Events: Go on pump

This is displayed along the top of the Learner Screen

5  Follow steps in Section 4.9 to set gas blender and Section 4.10 to go on pump

⑥ Press Save Step** button

** refers to the step having been modified and needs to be saved

⑦ In the Confirm Scenario File write dialog, select both Skills and Surgical Events

Hold <Ctrl> key to select multiple entries

⑧ Press Confirm button to save step changes to scenario file

### Figure 4.28 Step 3 edits

### Scenario File Step 4: Give Cardioplegia -

Follow these steps, refer to Figure 4.29

① Press Create New Step

② Type Give Cardioplegia for Scenario Step then press <Enter> on keyboard or click away with the mouse

③ Remove existing note in Skills field

④ Add following in Surgical Events: Give Cardioplegia

This is displayed along the top of the Learner Screen

5  Follow steps in Section 4.11 to deliver cardioplegia

⑥ Press Save Step** button

** refers to the step having been modified and needs to be saved

⑦ In the Confirm Scenario File write dialog, select both Cross Clamp and CP Delivery

Hold <Ctrl> key to select multiple entries

⑧ Press Confirm button to save step changes to scenario file

### Figure 4.29 Step 4 edits

### Scenario File Step 5: Event: Cannula malposition -

Follow these steps, refer to Figure 4.30

① Press Create New Step

② Type Event – Cannula malposition for Scenario Step then press <Enter> on keyboard or click away with the mouse

③ Remove existing note in Surgical Events field

4  Follow steps in Section 4.12 for Event: Cannula malposition

⑤ Press Save Step** button

** refers to the step having been modified and needs to be saved

⑥ In the Confirm Scenario File write dialog, select both Delivery Cannula 1 (%) and Surgical Events

Hold <Ctrl> key to select multiple entries

⑦ Press Confirm button to save step changes to scenario file

Figure 4.30 Step 5 edits

### Scenario File Step 6: Open Cross Clamp

Follow these steps, refer to Figure 4.31

① Press Create New Step

② Type Open Cross Clamp for Scenario Step then press <Enter> on keyboard or click away with the mouse

③ Add following note in Skills field: Return Delivery Cannula to 100% and open cross clamp

4  Follow steps in Section 4.13 to Open Cross Clamp

Also, set Delivery Cannula (%) to 100% since cannula malposition event has been fixed

⑤ Press Save Step** button

** refers to the step having been modified and needs to be saved

⑥ In the Confirm Scenario File write dialog, select Cross Clamp, Delivery Cannula 1 (%) and Skills

Hold <Ctrl> key to select multiple entries

⑦ Press Confirm button to save step changes to scenario file

### Figure 4.31. Step 6 edits

### Scenario File Step 7: Come off pump -

Follow these steps, refer to Figure 4.32

① Press Create New Step

② Type Come off pump for Scenario Step then press <Enter> on keyboard or click away with the mouse

③ Remove existing note in Skills field

4  Follow steps in Section 4.14 to come off pump

⑤ Press Save Step** button

** refers to the step having been modified and needs to be saved

⑥ In the Confirm Scenario File write dialog, select Skills

Hold <Ctrl> key to select multiple entries

⑦ Press Confirm button to save step changes to scenario file

Figure 4.32 Step 7 edits

#### 5

# Let’s go on ECMO! – an ECMO simulation

Approximately 60 minutes to complete this chapter.

This chapter steps through an ECMO simulation. You can follow along.

## SIM vs Califia Patient Module

Refer to Figure 5.1 , ①.

|  | Sections in this chapter headlined with this image refer to software-only simulation. No CPM or real pump is involved. |
| --- | --- |
|  | Sections in this chapter headlined with this image use an ECMO system connected to the CPM. |

## Use a Standardized Patient scenario

Select one of the six SP ECMO scenarios available, Figure 5.1, ②:

- SP VA ECMO Adult

- SP VA ECMO Infant

- SP VA ECMO Pediatric

- SP VV ECMO Adult

- SP VV ECMO Infant

- SP VV ECMO Pediatric

Press Continue button, ③, to step into the simulation.

Figure 5.1. Opening Screen "ECMO mode

## Patient Demographics

Refer to Figure 5.2, review and change the following:

① Age

② HCT

③ Gender

④ Weight

⑤ Height

### Figure 5.2. Patient demographics in Instructor's Panel

## CVP Gain

In Instructor’s Panel, press the Calibration Panel button under Califia Controls to bring up the Califia Calibration Panel, Figure5.3. Set CVP Gain as dictated in Table 5.1

| BSA (m2) | CVP Gain |
| --- | --- |
| < 1 | 3.5 |
| Between 1 & 2 | 3.0 |
| > 2 | 2.5 |

### Table 5.1. BSA vs CVP Gain

### Figure 5.3. Califia Calibration Panel > CVP Gain

## CVP

Set CVP to 10 mmHg

Fill Califia’s Reservoir until CVP is 10 mmHg, ① in Figure 5.4

### Figure 5.4. Fill CPM with water to match target CVP

Refer to Figure 5.5, go to Instructor’s Panel » Califia Controls

① Enter 10 for Target CVP then press <Enter>

② Volume in the virtual Califia’s Reservoir changes until CVP reaches 10, ③.

### Figure 5.5. Set Target CVP in SIM mode

## Arterial Pressure

Refer to Figure 5.6, in Instructor’s Panel » Cardiovascular Vars,

① Press Set Arterial Pressure button.

② Specify desired Systolic and Diastolic Arterial Pressures.

③ Press Target button.

④ It takes a few seconds to find values for LV Contractility and SVR.

⑤ Press Save Step to save all changes to scenario file.

Close Model Settings panel once targetting is complete.

### Figure 5.6. Set Systolic and Diastolic Arterial Pressure

## Ventilator Settings

While referring to Figure 5.7 , select Ventilator button in Instructor’s Panel and set values for:

① Rate

② FiO2

③ PIP

④ PEEP

### Figure 5.7. Ventilator settings

## Blood Gases

Refer to Figure 5.8, go to Instructor’s Panel » Blood Chemistry.

① Press Calculate Desired Gases button

② Specify desired pO2, pCO2 and pH for arterial blood

③ Press Target button

④ It takes several minutes to find values for HCO3, Dead Space and Pulmonary Shunt Fraction, the last two parameters are in the Pulmonary Vars section.

⑤ Press Save Step to save all changes to scenario file.

Close Model Settings panel once targetting is complete.

### Figure ‎5.8. Set Blood Gases

## Select Delivery and Drainage Cannulas

Refer to Figure 5.9  to select Delivery (arterial) and Drainage (venous) Cannulas. In Instructor’s Panel » Califia Controls,

① Press Cannula Editor button.

②a In Cannula Editor, select a Delivery Cannula from drop-down list then press Apply button.

③a Cannula label and corresponding valve opening is shown.

②b Change to Drainage Cannula, select cannula from drop-down list then press Apply button.

③b Cannula label and corresponding valve opening is shown.

Exit Cannula Editor

Pressure-Flow profiles for cannulas are modeled using respective electromechanical valves in CPM. The same performance profiles are also used in SIM mode.

### Figure 5.9 . Cannula selection

## Set Gas Blender

Refer to Figure 5.10, gas blender is set in the learner screen.

① Bring up bottom ribbon and select Oxygenator Controls button

② Set Gas Sweep Rate

③ Set Oxygenator FiO2 by pressing green button then turning blue knob

### Figure 5.10. Set Gas Blender

## Let’s go on pump!

Remove clamp on venous drainage line to drain patient while flowing using systemic pump.

Notice Arterial Pressure and CVP values decreasing as patient transitions support from native heart to heart lung machine, Figure 5.11.

### Figure 5.11. Arterial and CVP decreasing

Prime virtual circuit, refer to Figure 5.12

Bring up bottom ribbon

① Press Fluids Panel button

② Press HLM Controls button

In Fluids Panel

③a Select Crystalloid from Fluids drop-down menu

③b Enter 1000 mL of Volume

③c Light up the Priming button

③d Press Apply button

④ Notice Venous Reservoir fills up with this priming volume

### Figure 5.12. Prime circuit using Fluids Panel

To go on bypass, refer to Figure 5.13

② Start increasing systemic pump flow

Virtual systemic pump is a roller pump

③ Watch for arterial pressure waveform flattening though as flow is increased, mean pressure rises

### Figure 5.13. Starting bypass (SIM)

## Give Cardioplegia

Refer to Figure 5.14 , in Instructor’s Panel » Cardioplegia,

① Set slider to Ante (Antegrade Cardioplegia delivery)

② Close X Clamp

### Figure 5.14. Prepare for Cardioplegia delivery

Flow into Cardioplegia port. As flow is registered in the simulation software, the heart ECG will step through various patterns until the heart is completely arrested, Figure 5.15

### Figure 5.15 Heart ECG under Cardioplegia delivery

Refer to Figure 5.16

① Close X Clamp in HLM Controls panel

② Increase Cardioplegia pump flow

③ Notice ECG pattern change

### Figure 5.16. Heart ECG under Cardioplegia delivery (SIM)

## Event: Cannula malposition

Refer to Figure 5.17, in Instructor’s Panel » Califia Controls,

① Set Delivery Cannula (%) slider to about 25%

② Arterial line pressure increases

### Figure 5.17. Arterial Cannula set to 25% opened

Arterial line pressure should report a high value due to high resistance generated by the smaller arterial valve opening, Figure 5.18

[JN] – need line manometer photo

### Figure 5.18. Line pressure increase (line manometer photo)

Due to the high resistance generated by the smaller arterial valve opening, the arterial line pressure displayed using the pressure panel is very high,Figure 5.19

Line pressure (post-Oxygenator): 168 mmHg

Pre-Oxygenator pressure: 178 mmHg

### Figure 5.19. High arterial line pressures (SIM)

Set Delivery Cannula (%) back to 100% once the event is resolved.

## Remove Cross Clamp

Refer to Figure 5.20

① Open Cross Clamp in Instructor’s Panel » Cardioplegia

② Notice ECG pattern return to activity, ending in SR.

### Figure 5.20. Open Cross Clamp

Refer to Figure 5.21

① Open Cross Clamp in HLM Controls panel

② Notice ECG pattern return to activity, ending in SR.

### Figure 5.21. Open Cross Clamp (SIM)

## Come off pump!

Refer to Figure 5.22

Start weaning by partially occluding venous drain and filling patient.

① Notice patient’s heart starts to eject and venous pressure increases

Come-off systemic pump as patient’s arterial pressure returns to normal and CVP meets target value.

② Heart has taken over full body perfusion and CVP is normal value

### Figure 5.22. Wean off pump

Refer to Figure 5.23

① Start weaning by partially occluding venous drain and filling patient.

② Notice patient’s heart starts to eject and venous pressure increases

③ Come-off systemic pump as patient’s arterial pressure returns to normal and CVP meets target value.

④ Heart has taken over full body perfusion and CVP is normal value

### Figure 5.23. Wean off pump (SIM)

## Return to Opening Screen

Refer to Figure 5.24, press End Simulation button, confirm request to return to Opening Screen.

### Figure 5.24. End Simulation

## Create a CPB Simulation Scenario File

In this section, we will repeat the CPB case presented in this chapter however we will save all stages into a scenario file.

## 5.18 Access Scenario Editor

Refer to Figure 5.25, Press MENU button, ①, followed by Scenario Editor button, ②.

### Figure 5.25. Access MENU in Opening Screen then Scenario Editor

## 5.19 Save Scenario File

Refer to Figure 5.26, in Scenario Editor,

① Press File menu selection

② Select Save As….

③ Enter a new scenario file name such as Getting Started – CPB Scenario

This name is used to create a folder which will host the scenario file and ventilator.ini file

[Optional] – the Button Selector defaults to Other however a different button can be programmed to be associated with the new scenario file

④ Press the Create button to complete the new scenario file creation

Notice the Scenario Editor title reflects the new folder location and file name

⑤ Once new scenario file and folder created, return to File menu and press Close

Press Close button to close the MENU bar, this step is not shown.

### Figure 5.26. Save As... in Scenario Editor

## 5.20 New scenario file in Opening Screen

Refer to Figure 5.27, Opening Screen

① Location of Scenario File points to the newly created Scenario File

② The selected button is Other as specified in the Save As… dialog

③ Press Continue button to load new scenario file

### Figure 5.27. Opening Screen - New Scenario File

## 5.21 Scenario File Step 2: Select cannulas

The original scenario file was saved with patient’s blood gases configured as desired, section X.X. [Blood Gases in p. 25]

Next is to select cannulas. Follow these steps in Instructor’s Panel, refer to Figure 5.28

① Press Create New Step

② Step #2 is added to scenario table; type Select cannulas for Scenario Step then press <Enter> on keyboard or click away with the mouse

3 Follow steps in Section X.X to select cannulas

④ Press Save Step** button

** refers to the step having been modified and needs to be saved

⑤ In the Confirm Scenario File write dialog, select both Drainage Cannula and Delivery Cannula 1

Hold <Ctrl> key to select multiple entries

⑥ Press Confirm button to save step changes to scenario file

### Figure 5.28. Step 2 edits

## 5.22 Scenario File Step 3: Go on pump

Follow these steps, refer to Figure 5.29

① Press Create New Step

② Type Let’s go on pump for Scenario Step then press <Enter> on keyboard or click away with the mouse

③ Add following note in Skills field: Set gas blender and go on pump

④ Add following in Surgical Events: Go on pump

This is displayed along the top of the Learner Screen

5 Follow steps in Section X.X to set gas blender and Section X.X go on pump

⑥ Press Save Step** button

** refers to the step having been modified and needs to be saved

⑦ In the Confirm Scenario File write dialog, select both Skills and Surgical Events

Hold <Ctrl> key to select multiple entries

⑧ Press Confirm button to save step changes to scenario file

### Figure 5.29. Step 3 edits

## 5.23 Scenario File Step 4: Give Cardioplegia

Follow these steps, refer to Figure 5.30

① Press Create New Step

② Type Give Cardioplegia for Scenario Step then press <Enter> on keyboard or click away with the mouse

③ Add following note in Skills field: Set gas blender and go on pump

④ Add following in Surgical Events: Give Cardioplegia

This is displayed along the top of the Learner Screen

5 Follow steps in Section X.X to deliver cardioplegia

⑥ Press Save Step** button

** refers to the step having been modified and needs to be saved

⑦ In the Confirm Scenario File write dialog, select both Skills and Surgical Events

Hold <Ctrl> key to select multiple entries

⑧ Press Confirm button to save step changes to scenario file

### Figure 5.30. Step 4 edits

## 5.24 Scenario File Step 5: Event: Cannula malposition

Follow these steps, refer to Figure 5.31

① Press Create New Step

② Type Event – Cannula malposition for Scenario Step then press <Enter> on keyboard or click away with the mouse

③ Remove existing note in Surgical Events field

4 Follow steps in Section X.X for Event: Cannula malposition

⑤ Press Save Step** button

** refers to the step having been modified and needs to be saved

⑥ In the Confirm Scenario File write dialog, select both Delivery Cannula 1 (%) and Surgical Events

Hold <Ctrl> key to select multiple entries

⑦ Press Confirm button to save step changes to scenario file

### Figure 5.31. Step 5 edits

## 5.25 Scenario File Step 6: Open Cross Clamp

Follow these steps, refer to Figure 5.32

① Press Create New Step

② Type Open Cross Clamp for Scenario Step then press <Enter> on keyboard or click away with the mouse

③ Add following note in Skills field: Return Delivery Cannula to 100% and open cross clamp

4 Follow steps in Section X.X to Open Cross Clamp

Also, set Delivery Cannula (%) to 100% since cannula malposition event has been fixed

⑤ Press Save Step** button

** refers to the step having been modified and needs to be saved

⑥ In the Confirm Scenario File write dialog, select Cross Clamp, Delivery Cannula 1 (%) and Skills

Hold <Ctrl> key to select multiple entries

⑦ Press Confirm button to save step changes to scenario file

### Figure 5.32. Step 6 edits

## 5.26 Scenario File Step 7: Come off pump

Follow these steps, refer to Figure 5.33

① Press Create New Step

② Type Come off pump for Scenario Step then press <Enter> on keyboard or click away with the mouse

③ Remove existing note in Skills field

4 Follow steps in Section X.X to come off pump

⑤ Press Save Step** button

** refers to the step having been modified and needs to be saved

⑥ In the Confirm Scenario File write dialog, select Skills

Hold <Ctrl> key to select multiple entries

⑦ Press Confirm button to save step changes to scenario file

### Figure 5.33. Step 7 edits

Working on icons for document in process.

| Icon | Description | Icon | Image |
| --- | --- | --- | --- |
|  | Help / Note |  |  |
|  | Warning |  |  |
|  | Warning / Alert |  |  |
|  | Time / Duration |  |  |